## Slide 1: NSTX-U High-k Scattering receiver optical calibration
Xinhang Xu
7/25/2024

## Slide 2: introduction
3D structure of Receive Optical
Simplify of the NSTX-U High-k Scattering Receiver Geometry
Receiver optical calibration
Error analysis
Software development
Summary

## Slide 3
3D structure of Receive Optical
Z antenna Motor
R antenna Motor
Tilt antenna Motor
Toroidal antenna Motor
Window
Focus point

## Slide 4
Bay L Port Cover
Torus Center
Interaction Region (IR)
Window
RRW
zIR
RIR
xRW
φRA
ψRA
xRW = x-axis offset of receiver beam on window (nominally half of window width or 58 mm)
φRA = horizontal tilt of receiver array and optics (between 11.39 and 15.39°)
zIR = distance from window to interaction region (between 200 and 850 mm)
zIR = distance from window to interaction region (measured along plasma midplane)
RRW = major radius of vacuum window (fixed)
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
ψR = toroidal tilt angle between receiver beam and Radius vector (from torus center)
Toroidal tilt angle due to x-axis offset: ψX = tan‒1(xRW/RRW)
φRA = φRA + ψX
RRW2 = RRW2 + xRW2
RIR2 = RRW2 + zIR2 – 2RRW·zIR·cos(φRA)
ψRA = ψX + sin‒1((zIR/RIR)·sin(φRA))
ψR = ψRA + φRA
ψX
RRW
φRA
ψR
NSTX-U High-k Scattering Receiver Geometry(By Calvin and Xianzi)
Discussion: Here the xRW should be a variable parameter related to ,which is determined by relay mirror rotation angle. (Xinhang Xu,7/18/2024)

## Slide 5
1.Reference Point is used for relative calibration since we can use visible laser from receiver antenna to illuminate the receiver path and steer M2 until the laser point reach the center of the window. Then we can set that angle of M2 as =0, meaning the Reference point for M2 is home position
2.Absolute calibration
using visible laser and measure the angle of the window’s left edge as
using visible laser and measure the angle of the window’s
Cartesian coordinates where the window’s center is original point and x parallel to window while y vertical to window
Solving the equations, and (x,y)is the M2 axis position, where l is the half width of window.
Consider the ray along the window’s center =0), φRA (refer to page 3)should be:
when the M2’s rotation angle is ,the should be
High-k Receiver optical Calibration: find the position of M2 axis which determine the angle of
Bay L Port Cover
Reference Point
M2
M1
Torus Center
φRA
Distance from M2 axis to center of window :153.9mm
Right max angle of the edge:12.739degree
Left max angle of the edge:19.764degree
LOFF-AXIS=131.3mm

## Slide 6
Consider center of the window is (0,0),the Loff-axisis the distance from cw to axis offset ,than
Here (x,y) is the M2 axis position(page 2).
Window
xRW
φRA
M2
(x,y)
cw(0,0)
High-k Receiver optical Calibration: Calculate XRA

## Slide 7
Error analysis
1.1o error shift correspond to 4mm distance shift on the window, given the laser beam size is about 4mm, the error range could be ideally contained in 1o
2.This method only can make sure the incident angle error in 6o
x0=-58.65mm, y0=-154.42mm,
Given ,L=70mm

## Slide 8
Bay L Port Cover
Reference Point
Interaction Region (IR)
1.use visible laser from receiver antenna and steer M2 until the laser reach Launch Mirror point. record that angle of M2 as . The receiver antenna at the angle is possible exposed to the launch beam. Set the M2 angle below to avoid the possible exposing.
2. use visible laser from launch antenna and steer LM until the laser reach M2 center point. record that angle of LM as . The LM (Launch Mirror)at the angle possible illuminate the receiver antenna. Set the LM angle no equal to avoid the possible exposing.
Bay G Launch Mirror
(LM)
M2
M1
Way to Avoid exposing any of the receiver channels to the direct launch beam

## Slide 9
=atan(y/x)+2α
xRW= Loff-axis+(x-y·tan(ΦRA)) ψX = tan‒1(xRW/RRW)
φRA = φRA + ψX
RRW2 = RRW2 + xRW2
zIR=Lt*cos(θtilt)-Lw *cos(θtilt)- SR
RIR2 = RRW2 + zIR2 – 2RRW·zIR·cos(φRA)
ψRA = ψX + sin‒1((zIR/RIR)·sin(φRA))
ψR = ψRA + φRA
Z = γIR-Lt*sin(θtilt)
zIR=zIR/ cos(θtilt)
Torus Center
Interaction Region (IR)
Window
zIR
RIR
xRW
φRA
ψRA
ψX
RRW
φRA
ψR
RRW
Input parameters:
SR = receiver optical radical distance
θtilt = poloidal tilt angle of receiver optical
α = M2(mirror 2) rotation angle，α∈(-9o,6o)
γIR = IR height above midplane
Constant parameters :
Loff-axis = x-axis offset of receiver beam on window center (131.3mm).
RRW = major radius of vacuum window (fixed)1873.5mm.
Lt = optical total length from receiver antenna axis to focus point(2088mm).
Lw =length from receiver antenna axis to window1 (1132mm) when SR=0.
x ,y = M2(mirror 2) position(M2 center to Window center :154mm,angle:13.76o)
Depend parameters:
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
ψR = toroidal tilt angle between receiver beam and Radius vector (from torus center)
φRA = horizontal tilt of receiver array and optics (between 11.39 and 15.39°)
zIR = distance from window to interaction region (between 200 and 850 mm)
zIR = distance from window to interaction region (measured along plasma midplane)
Z = antenna height above midplane

## Slide 10
Control software

## Slide 11
Summary :
1.Analysis the method for absolute calibration of receiver angle, which ensure the error range in 60 .
2.Method to check the direct illumination angle between launch mirror and receiver mirror to avoid exposing
any of the receiver channels to the launch beam.
3. Develop software of receiver optical system for calculating the Interaction region and system position
control

## Slide 12
appendix

## Slide 13
Window center(0,0)
Mirror axis(x0,y0)
Right edge(L,0)
Left edge(-L,0)
Geometrical relation between angle and mirror axis