## Slide 1: NSTX-U High-k Scattering Calculation_2
Xinhang Xu
7/17/2024
reference:
https://app.box.com/s/qgvbzopdk8lqy0c985rz7479qrmi3uy3

## Slide 2
1.Define the optical geometry.
2. Get the constraint equation between the geometry parameters
3. High-k Receiver optical absolute calibration

## Slide 3
Bay L Port Cover
Torus Center
Interaction Region (IR)
Window
RRW
zIR
RIR
xRW
φRA
ψRA
xRW = x-axis offset of receiver beam on window (nominally half of window width or 58 mm)
φRA = horizontal tilt of receiver array and optics (between 11.39 and 15.39°)
zIR = distance from window to interaction region (between 200 and 850 mm)
zIR = distance from window to interaction region (measured along plasma midplane)
RRW = major radius of vacuum window (fixed)
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
ψR = toroidal tilt angle between receiver beam and Radius vector (from torus center)
Toroidal tilt angle due to x-axis offset: ψX = tan‒1(xRW/RRW)
φRA = φRA + ψX
RRW2 = RRW2 + xRW2
RIR2 = RRW2 + zIR2 – 2RRW·zIR·cos(φRA)
ψRA = ψX + sin‒1((zIR/RIR)·sin(φRA))
ψR = ψRA + φRA
ψX
RRW
φRA
ψR
NSTX-U High-k Scattering Receiver Geometry(By Calvin)
Discussion: Here the xRW should be a variable parameter related to ,which is determined by relay mirror rotation angle. (Xinhang Xu,7/18/2024)

## Slide 4
1.Reference Point is used for relative calibration since we can use visible laser from receiver antenna to illuminate the receiver path and steer M2 until the laser point reach the center of the window. Then we can set that angle of M2 as =0, meaning the Reference point for M2 is home position
2.Absolute calibration
using visible laser and measure the angle of the window’s left edge as
using visible laser and measure the angle of the window’s right edge as
Set up the Cartesian coordinates where the window’s center is original point and x parallel to window while y vertical to window
Solving the equations, and (x,y)is the M2 axis position, where l is the width of window.
Consider the ray along the window’s center =0), φRA (refer to page 3)should be:
when the M2’s rotation angle is ,the should be
High-k Receiver optical Calibration: find the position of M2 axis which determine the angle of
Bay L Port Cover
Reference Point
M2
M1
Torus Center
φRA
Dm2 to axis of window :153.9mm
Angle :
Right max angle of the edge:12.739degree
Left max angle of the edge:19.764degree
LOFF-AXIS=131.3mm

## Slide 5
Consider center of the window is (0,0),the Loff-axisis the distance from cw to axis offset ,
Here (x,y) is the M2 axis position(page 2).
Window
xRW
φRA
M2
(x,y)
cw(0,0)
High-k Receiver optical Calibration: Calculate XRA

## Slide 6
Bay L Port Cover
Reference Point
Interaction Region (IR)
1.use visible laser from receiver antenna and steer M2 until the laser reach Launch Mirror point. record that angle of M2 as . The receiver antenna at the angle is possible exposed to the launch beam. Set the M2 angle below to avoid the possible exposing.
2. use visible laser from launch antenna and steer LM until the laser reach M2 center point. record that angle of LM as . The LM (Launch Mirror)at the angle possible illuminate the receiver antenna. Set the LM angle no equal to avoid the possible exposing.
Bay G Launch Mirror
(LM)
M2
M1
Way to Avoid exposing any of the receiver channels to the direct launch beam

## Slide 7
=atan(y/x)+2α
xRW= Loff-axis-(x-y·tan(ΦRA)) ψX = tan‒1(xRW/RRW)
φRA = φ= tan‒1(xRW/RRWRA )+ ψX
RRW2 = RRW2 + xRW2
RIR2 = RRW2 + zIR2 – 2RRW·zIR·cos(φRA)
ψRA = ψX + sin‒1((zIR/RIR)·sin(φRA))
ψR = ψRA + φRA
z = Z+Lt*sin(θtilt)
Torus Center
Interaction Region (IR)
Window
zIR
RIR
xRW
φRA
ψRA
ψX
RRW
φRA
ψR
RRW
Finally ,we have
For receiver carriage with four axis ,we can arrange the focus distance :ZIR,the toroidal angle α ，the poloidal tilt angle : and the antenna height Z .input(ZIR, α, ,Z) we will get the IR position(ψRA , RIR,z) and receiver angle ψR
xRW = x-axis offset of receiver beam on window (nominally half of window width or 58 mm)
φRA = horizontal tilt of receiver array and optics (between 11.39 and 15.39°)
zIR = distance from window to interaction region (between 200 and 850 mm)
zIR = distance from window to interaction region (measured along plasma midplane)
RRW = major radius of vacuum window (fixed)1873.5mm
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
ψR = toroidal tilt angle between receiver beam and Radius vector (from torus center)
z = IR height above midplane
Z = antenna height above midplane
Lt = receiver optical length
θtilt = poloidal tilt angle of receiver optical
x ,y = M2(mirror 2) position
α = M2(mirror 2) rotation angle

## Slide 8
alignment procedure:
Take down the cylindrical lens and meniscus lens .
mount the visible laser on the central of the antenna and adjust the visible laser angle to reach to the center of the convex lens.
Take down all the receiver lens and check if the laser point is at the center of the window.
If the laser not on the center of the window ,adjust the mirror until the laser point reach at the center of the window.
Remount the receiver lens after alignment .

## Slide 9
=atan(y/x)+2α
xRW= Loff-axis-(x-y·tan(ΦRA)) ψX = tan‒1(xRW/RRW)
φRA = φRA + ψX
RRW2 = RRW2 + xRW2
RIR2 = RRW2 + zIR2 – 2RRW·zIR·cos(φRA)
ψRA = ψX + sin‒1((zIR/RIR)·sin(φRA))
ψR = ψRA + φRA
Z = γIR-Lt*sin(θtilt)
zIR=Lt*cos(θtilt)-Lw-SR
zIR=zIR/ cos(θtilt)
Torus Center
Interaction Region (IR)
Window
zIR
RIR
xRW
φRA
ψRA
ψX
RRW
φRA
ψR
RRW
Constant parameters :
xRW = x-axis offset of receiver beam on window center (131.3mm).
RRW = major radius of vacuum window (fixed)1873.5mm.
Lt = optical total length from receiver antenna axis to focus point(2088mm).
Lw =length from receiver antenna axis to window1 (1132mm) when SR=0.
x ,y = M2(mirror 2) position(M2 center to Window center :154mm,angle:13.76o)
Depend parameters:
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
ψR = toroidal tilt angle between receiver beam and Radius vector (from torus center)
φRA = horizontal tilt of receiver array and optics (between 11.39 and 15.39°)
zIR = distance from window to interaction region (between 200 and 850 mm)
zIR = distance from window to interaction region (measured along plasma midplane)
Z = antenna height above midplane
Input parameters:
SR = receiver optical radical distance
θtilt = poloidal tilt angle of receiver optical
α = M2(mirror 2) rotation angle
γIR = IR height above midplane

## Slide 10
Background Information Slides on scattering wavenumber and Pitch Angle Effects
By Calvin

## Slide 11
Bay G Launch Mirror
Torus Center
Window
NSTX-U High-k Scattering Launch Geometry
Interaction Region (IR)
ψL
RLM
zLM
RIR
RRW
ψRA
RRW = major radius of vacuum window
RIR = major radius of interaction region
ψRA = toroidal tilt angle of interaction region (w.r.t. plane of vacuum window)
RLM = major radius of launch mirror (fixed)
ψLM = toroidal tilt angle of launch mirror (w.r.t. plane of vacuum window)
zLM = distance from launch steering mirror to interaction region
ψL = toroidal tilt angle between launch beam and Radius vector (from torus center)
ψLM
zLM2 = RLM2 + RIR2 – 2RLM·RIR·cos(ψLM - ψRA)
ψL = sin‒1((RLM/zLM)·sin(ψLM - ψRA))

## Slide 12
NSTX-U High-k Scattering Wavenumber Calculations (Ignore Magnetic Pitch Factor)
Window
If the launch and receive beams are toroidally aligned, then the radial component is zero. A slight toroidal tilt is, however, required to minimize unscattered power illuminating the sensitive receiver mixers.
Calculate the amount of radial wavenumber required to toroidally steer the scattered beams.
Then calculate the kθ components due to the vertical tilt angle θ of a given channel.
ψR
Interaction Region (IR)
ψL
Radial Vector
Launch Beam
Receive Beam
Poloidal View of High-k Receiver
kR = k0·(cosψR – cosψL)
kθ = k0·sinθ

## Slide 13
Effects of Magnetic Pitch Angle
The magnetic pitch angle α is the angle between the total magnetic field BTOT and the toroidal magnetic field BTOR due to the presence of a non-zero poloidal magnetic field Bθ.
The magnetic pitch angle causes the poloidal and toroidal vectors to rotate by the factor α. Note that this is much larger on a spherical tokamak than on a conventional tokamak.
This modifies the calculated density fluctuation kθS and kRS components that scatter from the launch beam towards the receiver.
This increases the kθ component for downwards scattering and decreases the kθ component for upwards scattering, as well as adding a poloidal scattering angle θ term to the radial scattering kR component.
kRS = k0·[cosψR·cosθS – cosψL]
kθS = k0·[(sinψR·cosθS – sinψL)·sinα + sinθS·cosα]
α
Toroidal Axis
Poloidal Axis
θS
Launch Beam
Scattered Beam

## Slide 14
NSTX-U Scattering Calculator v7: Instructions on Use
There are 2 plasma parameters that can be optionally set by the user:
Cell E4: Plasma major radius R0 , set between 800 and 1020 mm
Cell E5: Plasma minor radius a, set between 495 and 650 mm
There are 4 parameters that can be set by the user:
Cell C11: Distance from receiver window (RW) to interaction region (IR) zIR, set between 200 and 850 mm
Cell C12: Receiver array horizontal tilt φRA, set between 10.39° and 16.39°
Cell C13: Receiver array vertical tilt θRA, set between -5.0° and +5.0°
Cell C15: Magnetic pitch angle α, input from the Pitch Angle tab using the calculated IR major radius (cell M14)
The user can set the 2 plasma parameters, although they don’t affect any of the spreadsheet calculations except for calculating the approximate value of r/a (cell M15).
The user enters a value for the θRA to focus on upwards or downwards scattering. Note that deepest into the plasma (zIR = 850 mm), θRA has to set to 0.0°, as beams must pass through the vacuum window. As zIR is reduced, the vertical tilt can be increased. As per the 5/11/2022 optics design, the tilt is limited to:
|θRA| ≤ tan−1(138.66/zIR) − 9.26°, i.e. a 1° tilt is possible up to 766 mm, a 2° tilt is possible up to 696 mm, a 3° tilt is possible up to 638 mm, a 4° tilt is possible up to 588 mm, and a 5° tilt is possible up to 545 mm
The user enters values for zIR and φRA to put the interaction region at a particular desired major radius RIR (cell M14). The user then enters the appropriate value for α corresponding to this value of RIR.
Finally, the user can try to optimize the scattering performance (minimize radial wavenumbers) by adjusting the horizontal tilt angle φRA.

## Slide 15
Background Information Slides on Pitch Angle Effects

## Slide 16: Effects of Magnetic Pitch Angle (by Xianzi Liu)
Mid Plane
Torus Center
IR

## Slide 17
Build a cartesian coordinates (x,r,z) with the origin at IR (Interaction Region)
: major radius direction, from the torus center to IR. Radial direction.
: perpendicular to the mid plane, upward. The traditional poloidal direction.
: perpendicular to and , (x,r,z) form a right-hand system. The traditional toroidal direction.
IR
radial
poloidal

## Slide 18
Build another cartesian coordinates (x’,r,z’) with the origin at IR
Compared to the previous one, this one rotates a degree of (pitch angle) around axis .
: major radius direction, from the torus center to IR. Radial direction.
: parallel to the total magnetic field. The newly-defined toroidal direction.
: perpendicular to and , (x’,r,z’) form a right-hand system. The newly-define poloidal direction.
radial
poloidal

## Slide 19
The IR (Interaction Region) is on the mid plane, we can treat that the radial direction (the connection between torus center and IR) is perpendicular to the magnetic flux surface.
But if the IR is above or below the mid plane, the radial direction is not strictly perpendicular to the magnetic flux surface, then maybe we should treat it differently if the deviation is

## Slide 21
Dantenna2axis=149.85mm

## Slide 22
Danntena2lenB1=637.84mm

## Slide 23
Lens B1 to Focus:79+40+70+155+16.35+940=1300.3mm
Optical length = Lens B1 to Focus+Danntena2lenB1+ Dantenna2axis=1300.3+149.85+637.84=2088mm
Lw= Optical length –W12Fo=2088-940-16.35=1132mm